## Octant 1
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 1 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
RETEST AFTER CLIPS FITTED 16/11/2006
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
1 | B6 P1 | US15 | 05-14/1 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
6 | B6 P4 | US15 | 12-15/1 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
8 | B6 P5 | US15 | 12-15/1 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
2 | B6 P2 | US15 | 12-15/1 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
4 | B6 P3 | US15 | 12-15/1 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
10 | B4 P5 | US15 | 23-15/2 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
3 | B5 P2 | US15 | 23-15/3 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
5 | B5 P3 | US15 | 23-15/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | fail | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
7 | B5 P4 | US15 | 23-15/3 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
9 | B6 P6 | US15 | 23-15/4 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
11 | B6 P7 | US15 | 23-15/4 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
12 | B5 P1 | US15 | 05-14/2 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
13 | B4 P1 | US15 | 05-14/3 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | fail
14 | B3 P1 | US15 | 05-14/4 | 0 | 6 | ok | 1 | 5 | fail | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | fail
16 | B3 P2 | US15 | 23-15/1 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | fail
18 | B3 P3 | US15 | 23-15/1 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
21 | B3 P4 | US15 | 23-15/1 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
15 | B4 P2 | US15 | 23-15/2 | 0 | -6 | fail | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | fail | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | fail | 10 | 5 | ok | 11 | 6 | ok
17 | B4 P3 | US15 | 23-15/2 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
19 | B4 P4 | US15 | 23-15/2 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
20 | B5 P5 | US15 | 23-15/3 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | fail
22 | B5 P6 | US15 | 24-15/3 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | fail
Note on G12: CableTest: HV backplane R 53k
Note on G15: CableTest: 40.5k!!!
Note on G25: CableTest: SELECT-VddRet - 3.94k
Note on J21: CableTest: 21.11.06 Temp1-VddRet - 99.0E+36 retest after replugging PPB1 - the same
Note on M8: CableTest: 100k ish
Note on M9: CableTest: 40.8k
Note on M10: CableTest: HV backplane R 90k static
Note on M11: CableTest: Backplane R varies between 80 and 100k, not static???
Note on M13: CableTest: High R on HVRET 11k. Present also on 11/11/2006, but missed
Note on M21: CableTest: HVBias-HVRet - 75.8k
Note on M22: CableTest: HVBias-HVRet - 44.7k
Note on P11: CableTest: 50 ish k static
Note on P12: CableTest: HV Backplane 82k
Note on P16: CableTest: 66k static
Note on P25: CableTest: HVBias-HVRet - 104k
Note on S11: CableTest: Strictly fail, SELECT is right on the cut at 5k. Should be OK.
Note on S12: CableTest: HV backplane R 41k
Note on S15: CableTest: 41.3k
Note on V8: CableTest: 45k
Note on V11: CableTest: Select 4.93k.
Note on V12: CableTest: HV backplane R 100k after two training cycles
Note on V16: CableTest: circa 53k
Note on V21: CableTest: HVBias-HVRet- 42.1k
Note on V26: CableTest: HVBias-HVRet -57.6k
Note on Y8: CableTest: OK AFTER RETEST WITH NEW HV PATCH CARD
Note on Y9: CableTest: SELECT 4.99k
Note on Y13: CableTest: HV backplane R 60k static
Note on Y15: CableTest: 50k
Note on Y21: CableTest: HVBias-HVRet - 42.9k
Note on AB9: CableTest: 55k after two goes at training
Note on AB11: CableTest: 43k
Note on AB14: CableTest: HV Backplane R static at 80k
Note on AB21: CableTest: HVBias-HVRet - 40.1k
Note on AB26: CableTest: HVBias-HVRet 88.1k
Note on AE6: CableTest: After training once, 39k
Note on AE20: CableTest: Hvret - 40.8k
Note on AE22: CableTest: HVBias-HVRet - 58.5k
Note on AH21: CableTest: HVBias-HVRet - 53.3k
Note on AH25: Peter W Phillips: Only one working vcsel (known from Oxford) 21.11.06 VVCSEL-VddRet - 137k
Note on AK16: CableTest: circa 55k
Note on AK22: Peter W Phillips: SR1: VddRetSense linked to VddRet at PPB1 due to LMT fault 21.11.06 HVBias-HVRet - 40.8k
Note on AN14: CableTest: 40.3k!!!
Note on AN15: CableTest: 40.3k
Note on AN20: CableTest: SELECT-VddRet - 3.96k
Note on AN21: CableTest: VVCSEL-VddRet - 137k HVRet-VddRet - 99.0E+36 HVBias-HVRet - 2.1M 21.11.06 retest after replugging - the same
Note on AN22: CableTest: HVBias-HVRet - 101k
Note on AN28: CableTest: SELECT-VddRet - 4.84k
Note on AN29: CableTest: SELECT-VddRet - 4.74k

## 0ctant 2
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 2 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
RETEST AFTER CLIPS FITTED 16/11/2006. Some rows of this table had incorrect +/-Z annotations, corrected 16/11/2006
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
23 | B6 P8 | US15 | 23-15/4 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
25 | B6 P9 | US15 | 23-15/4 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | fail | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
24 | B4 P6 | US15 | 24-15/2 | 0 | -6 | Fixed | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 |  | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
27 | B5 P9 | US15 | 24-15/3 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
26 | B6 P10 | US15 | 24-15/4 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
28 | B6 P11 | US15 | 24-15/4 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
30 | B6 P12 | US15 | 24-15/4 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
32 | B6 P13 | US15 | 24-15/4 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
| B6 P14 | US15 | 27-15/1 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | fixed | 11 | -6 | ok
31 | B5 P11 | US15 | 27-15/1 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | fail | HVBias-HVRet
29 | B5 P10 | US15 | 27-15/1 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
35 | B3 P5 | US15 | 23-15/1 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
38 | B3 P6 | US15 | 24-15/1 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
40 | B3 P7 | US15 | 24-15/1 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
42 | B3 P8 | US15 | 24-15/1 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
43 | B4 P10 | US15 | 24-15/1 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
34 | B5 P7 | US15 | 24-15/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | fail | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
36 | B5 P8 | US15 | 24-15/3 | 24 | -6 | fail | 25 | -5 | fail | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
37 | B4 P7 | US15 | 24-15/2 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | fail | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
39 | B4 P8 | US15 | 24-15/2 | 24 | -6 | fail | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
41 | B4 P9 | US15 | 24-15/2 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
44 | B5 P12 | US15 | 27-15/1 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | fail | 21 | -4 | fail | 22 | -5 | ok | 23 | -6 | ok
Note on D26: CableTest: it was 25-15/2
Note on D27: CableTest: it was 25-15/2
Note on D28: CableTest: it was 25-15/2
Note on D29: CableTest: crate number was wrong (it was 27-15/3) the correct one is crate number 1 (27-15/1)
Note on G6: CableTest: 60k static
Note on G8: CableTest: VccRSen-VccRet and VddRSen-VddRet high impedance. LMT failure, "fixed" by soldering at PPB1 (PWP). Values now well between limits.
Note on G9: CableTest: Ground card was not properly in place for first test
Note on G25: Peter W Phillips: SR1: Low R Select-VddReturn 21.11.06 SELECT-VddRet - 2.67k
Note on G27: CableTest: SELECT-VddRet - 917Ohms
Note on J19: CableTest: HVBias-HVRet - 57.2k
Note on J25: CableTest: 21.11.06 SELECT-VddRet - 2.67k
Note on J28: CableTest: HVBias-HVRet - 40.4k
Note on M7: CableTest: Hvbias now 62k (was 108kOhm last week?)
Note on M15: CableTest: VccRet open circuit-retest 15/11 still vccret fault, then reconnected and PASS!
Note on M22: Cable test: HVBias-HVRet - 41.2k
Note on S6: Peter W Phillips: Only one wokring vcsel (known from Oxford), apparently not due to LMT break
Note on S12: CableTest: 64k
Note on S13: CableTest: 64k
Note on S28: CableTest: HVBias-HVRet - 46.2k
Note on V7: CableTest: 57k static
Note on V15: CableTest: HV Backplane resistant settles to 56k 51k on retest
Note on V19: CableTest: HVBias-HVRet- 52.9k
Note on Y24: Peter W Phillips: SR1: slioghtly high R in Vcc loop
Note on Y26: CableTest: VVCSEL-VddRet - 8.32k
Note on AB7: CableTest: VVCSEL-VddRet 136K (missing vcsel known from Oxford)
Note on AB16: CableTest: failed first test due to faulty HV passthrough card, but is OK
Note on AB22: CableTest: HVBias-HVRet - 44.9k
Note on AB24: CableTest: Vcc-VccSense - 176Ohms
Note on AE8: CableTest: Hvret to VddRet short - suspect deliberate act at PPB1? Retested in the afternoon but still very resistance 12.3 Ohms and we'll test with dPPB1
Note on AE9: CableTest: HV backplane resistance 56k after 2 goes
Note on AE29: CableTest: SELECT-VddRet - 4.81k
Note on AH20: Peter W Phillips: SR1: VddRetSense linked to VddRet at PPB1 due to LMT fault.
Note on AH22: CableTest: HVBias-HVRet - 47.4k
Note on AH29: CableTest: SELECT-VddRet - 4.87k
Note on AK7: CableTest: HV backplane R 45k
Note on AK14: CableTest: VccRet-VccRetSen 10.3k broken wire fixed by Jacek. OK
Note on AN6: CableTest: SELECT-VddRet right on the lower limit
Note on AN7: CableTest: Hvbias 149k on first test. Dropped below 40k after cable wiggle and retest. 16th November test again - passes striaght away!
Note on AN15: CableTest: VVCSEL-VddRet 136kOhm (new missing vcsel)
Note on AN22: CableTest: HVBias-HVRet -76k
Note on AN24: Peter W Phillips: Vcsel bias repaired in SR1

## Octant 3
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 3 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
3 | B5 P14 | USA15 | 22-14/1 | 0 | -6 | ok | 1 | -5 | FIXED | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
5 | B5 P15 | USA15 | 22-14/1 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
7 | B5 P16 | USA15 | 22-14/1 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
6 | B6 P18 | USA15 | 22-14/2 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
8 | B6 P19 | USA15 | 22-14/2 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
9 | B6 P20 | USA15 | 22-14/2 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
11 | B6 P21 | USA15 | 22-14/2 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
1 | B6 P15 | USA15 | 22-14/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
2 | B6 P16 | USA15 | 22-14/3 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | FAIL | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
4 | B6 P17 | USA15 | 22-14/3 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | FAIL | 47 | 6 | ok
10 | B4 P15 | USA15 | 26-11/2 | 12 | 6 | ok | 13 | 5 | FAIL | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | FAIL
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
20 | B5 P17 | USA15 | 22-14/1 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
12 | B5 P13 | USA15 | 22-14/3 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
21 | B3 P12 | USA15 | 26-11/1 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
18 | B3 P11 | USA15 | 26-11/1 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
16 | B3 P10 | USA15 | 26-11/1 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | FAIL
19 | B4 P14 | USA15 | 26-11/2 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
22 | B5 P18 | USA15 | 26-11/3 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | FAIL | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
17 | B4 P13 | USA15 | 26-11/4 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | FAIL | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
15 | B4 P12 | USA15 | 26-11/4 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
13 | B4 P11 | USA15 | 26-11/4 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
14 | B3 P9 | USA15 | 26-11/4 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
Note on G6: CableTest: Hvbias-Hvret 86kOhm
Note on G22: CableTest: HVBias-HVRet 51.9k
Note on G23: CableTest: HVBias-HVRet 41.7k
Note on J6: CableTest: Hvret open? (high resistance to Vddret and Hvbias). Hvbias was fixed by intervention of P.Phillips. Hvret to VDDRET now has TOO LOW value.[PWP - expected - HVRET now connecetd to VCCRET on PPB1. Best we can do.] Corallary - this was seen during test at SR1!
Note on J16: CableTest: VPIN-VddRet -3.17V
Note on J22: CableTest: HVBias-HVRet 47.7k
Note on J26: CableTest: Vdd-VddSense 43.0k
Note on J29: CableTest: HVBias-HVRet 48.0k
Note on M10: CableTest: HVBias-Hvret 41.2kO
Note on M14: CableTest: Suspect vccsel, VVCSEL-VddRet is 136 kOhm (new)
Note on M21: Cable test: HVBias-HVRet 44.4k
Note on M24: CableTest: HVBias-HVRet 66.0k
Note on M27: CableTest: HVBias-HVRet 60.0k
Note on P24: CableTest: HVBias-HVRet 40.6k
Note on P25: CableTest: HVBias-HVRet 44.4k
Note on P26: CableTest: HVBias-HVRet 109k
Note on P29: CableTest: HVBias-HVRet 43.2k
Note on S10: CableTest: HVBias-Hvret 44.8kO
Note on S23: CableTest: HVBias-HVRet 42.2k
Note on S24: CableTest: HVBias-HVRet 43.8k
Note on S26: CableTest: HVBias-HVRet 64.8k
Note on V21: CableTest: HVBias-HVRet 54.0k
Note on V23: CableTest: HVBias-HVRet 52.1k
Note on Y14: CableTest: Hvbias 77kOhm
Note on AB25: CableTest: Vdd-VddSense 10.5k
Note on AE9: CableTest: it was Mohms between HVBias-Hvret but then HV feed-through board replugged and it PASS
Note on AE23: CableTest: HVBias-HVRet 46.9k
Note on AE29: CableTest: HVBias-HVRet 63.8k
Note on AK15: CableTest: CableTest: Suspect vccsel, VVCSEL-VddRet is 136 kOhm (known from Oxford)
Note on AK22: CableTest: HVBias-HVRet 48.8k
Note on AK23: CableTest: HVBias-HVRet 45.2k
Note on AK29: CableTest: HVBias-HVRet 40.2k
Note on AN22: CableTest: HVBias-HVRet 42.2k
Note on AN23: CableTest: Temp1-VddRet 99.0E+36 VPIN-VddRet -3.17 VVCSEL-VddRet 137k

## Octant 4
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 4 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
24 | B4 P16 | USA15 | 26-11/2 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | fail | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | fail | 10 | -5 | ok | 11 | -6 | ok
25 | B6 P23 | USA15 | 26-11/2 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
23 | B6 P22 | USA15 | 26-11/3 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
26 | B6 P24 | USA15 | 28-11/3 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
28 | B6 P25 | USA15 | 28-11/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
27 | B5 P21 | USA15 | 28-11/3 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
30 | B6 P26 | USA15 | 28-11/4 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
32 | B6 P27 | USA15 | 28-11/4 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
33 | B6 P28 | USA15 | 28-11/4 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | fail | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
29 | B5 P22 | USA15 | 29-11/1 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
31 | B5 P23 | USA15 | 29-11/1 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
35 | B3 P13 | USA15 | 26-11/1 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | fail | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
36 | B5 P20 | USA15 | 26-11/3 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
34 | B5 P19 | USA15 | 26-11/3 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
37 | B4 P17 | USA15 | 28-11/3 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
39 | B4 P18 | USA15 | 28-11/4 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
41 | B4 P19 | USA15 | 29-11/1 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
43 | B4 P20 | USA15 | 29-11/1 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
38 | B3 P14 | USA15 | 29-11/3 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | fail | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
40 | B3 P15 | USA15 | 29-11/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | FAIL
42 | B3 P16 | USA15 | 29-11/3 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | Fail | 28 | -2 | fail | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
44 | B5 P24 | USA15 | 29-11/3 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
Note on G8: CableTest: 40.6k
Note on G9: Peter W Phillips: link 1 dead at combined test, but vcsel connecdtivity was OK.
Note on G11: Peter W Phillips: Only one working vcsel (known from Oxford) 20.11.06 Hvbias-Hvret 53.0k
Note on G14: CableTest: Hvbias-Hvret 69.8k
Note on J14: CableTest: Hvbias-Hvret 47.2k
Note on J20: CableTest: HVBias-HVRet 41.7k
Note on M13: CableTest: Hvbias-Hvret 83.4k training - 80.7k
Note on M20: CableTest: HVBias-HVRet 62.7k
Note on M27: CableTest: HVBias-HVRet 41K
Note on P16: CableTest: HVBias-HVRet 45.5KOhms
Note on P24: CableTest: HVBias-HVRet 41k
Note on P28: CableTest: VVCSEL-VddRet 137KOhms Hvbias 45.4K
Note on S23: CableTest: HVBias-HVRet 55.1KOhms
Note on S28: CableTest: HVBias-HVRet 88.4KOhms
Note on V26: CableTest: VVCSEL-VddRet 8.3 kOhms - very low reading
Note on Y6: CableTest: fail:selectVddRet=4.94K
Note on Y14: Peter W Phillips: Only one working vcsel (known from Oxford)
Note on Y21: CableTest: HVBias-HVRet 51.7k
Note on Y24: CableTest: HVBias-HVRet 50.2kOhms
Note on Y27: CableTest: HVBias-HVRet 40.8K
Note on AB14: CableTest: Hvret shorted to Vddret -18.9Ohms Hvbias- Hvret -51.8k
Note on AB19: CableTest: Temp1-VddRet 15.6k
Note on AB20: CableTest: HVBias-HVRet 41.3k
Note on AE24: CableTest: HVBias-HVRet 41.5 kOhms
Note on AE27: CableTest: HVBias-HVRet 40.0K gives FAIL in the data
Note on AH6: CableTest: HVRet-VddRet=18.9. tested twice. Swapped with ch.8 and then ch8 had the same failure.
Note on AH14: CableTest: Hvbias -40.2k
Note on AN21: Peter W Phillips Only one working vcsel (known from Oxford)
Note on AN27: CableTest: HVRet-VddRet 99.0E+36 ohms HVBias-HVRet 2.10M Ohms

## Octant 5
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 5 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
9 | B6 P34 | USA15 | 28-14/1 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
11 | B6 P35 | USA15 | 28-14/1 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
6 | B6 P32 | USA15 | 28-14/3 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
10 | B4 P25 | USA15 | 28-14/3 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
8 | B6 P33 | USA15 | 28-14/4 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
5 | B5 P27 | USA15 | 28-14/4 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
7 | B5 P28 | USA15 | 28-14/4 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
2 | B6 P30 | USA15 | 29-14/1 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
4 | B6 P31 | USA15 | 29-14/1 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
3 | B5 P26 | USA15 | 29-14/1 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
1 | B6 P29 | USA15 | 29-14/2 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
22 | B5 P30 | USA15 | 26-14/1 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
17 | B4 P23 | USA15 | 28-14/2 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
16 | B3 P18 | USA15 | 28-14/2 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
21 | B3 P20 | USA15 | 28-14/2 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
18 | B3 P19 | USA15 | 28-14/3 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
19 | B4 P24 | USA15 | 28-14/3 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
20 | B5 P29 | USA15 | 28-14/4 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
15 | B4 P22 | USA15 | 29-14/1 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
12 | B5 P25 | USA15 | 29-14/2 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
13 | B4 P21 | USA15 | 29-14/2 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
14 | B3 P17 | USA15 | 29-14/2 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
Note on B29: Peter W Phillips: "The Row of Woe"!
Note on J13: Peter W Phillips: Only one working vcsel (known from Oxford)
Note on Y29: Peter W Phillips: SR1: HVRET linked to VccRetSense at PPB1 due to LMT fault. Also VddRet linked to VddRetSense and Vdd to VddSense at PPB1 (precuationary due to tape damage).
Note on AB29: Peter W Phillips: VddRetSense linked to VddRet at PPB1. VccSense linked to Vcc at PPB1.
Note on AN29: Peter W Phillips: Sr1: VccSense linked to Vcc at PPB1 due to fault. Also VddSense linked to Vdd (precuationary due to damaged tape)

## Octant 6
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 6 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
24 | B4 P26 | USA15 | 25-11/2 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | -1 |  | 7 | -2 |  | 8 | -3 |  | 9 | -4 |  | 10 | -5 |  | 11 | -6
29 | B5 P34 | USA15 | 25-11/3 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
31 | B5 P35 | USA15 | 25-11/3 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
33 | B6 P42 | USA15 | 25-11/3 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
26 | B6 P38 | USA15 | 25-11/4 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
28 | B6 P39 | USA15 | 25-11/4 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
30 | B6 P40 | USA15 | 25-11/4 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
32 | B6 P41 | USA15 | 25-11/4 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
27 | B5 P33 | USA15 | 26-14/1 | 0 | 6 |  | 1 | 5 |  | 2 | 4 |  | 3 | 3 |  | 4 | 2 |  | 5 | 1 |  | 6 | -1 |  | 7 | -2 |  | 8 | -3 |  | 9 | -4 |  | 10 | -5 |  | 11 | -6
23 | B6 P36 | USA15 | 28-14/1 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
25 | B6 P37 | USA15 | 28-14/1 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
38 | B3 P22 | USA15 | 25-11/1 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
40 | B3 P23 | USA15 | 25-11/1 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
42 | B3 P24 | USA15 | 25-11/1 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
43 | B4 P30 | USA15 | 25-11/1 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
37 | B4 P27 | USA15 | 25-11/2 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
39 | B4 P28 | USA15 | 25-11/2 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
41 | B4 P29 | USA15 | 25-11/2 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
44 | B5 P36 | USA15 | 25-11/3 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
36 | B5 P32 | USA15 | 26-14/1 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
34 | B5 P31 | USA15 | 26-14/1 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
35 | B3 P21 | USA15 | 28-14/2 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
Note on G11: Peter W Phillips: link 0 dead (known from SR1). Ileda has good connection, but other tape fully laminated, so cannot be probed.
Note on G21: Peter W Phillips: Only one working vcsel (known from SR1)
Note on G28: Peter W Phillips: Zero IPIN during SR1 test, however diode test OK.
Note on P25: Peter W Phillips: Only one working vcsel (known from Oxford)
Note on V12: Peter W Phillips: Oxford: VddSense-Vdd connected at PPB1
Note on V28: Peter W Phillips: Zero IPIN during SR1 test, however diode test OK.
Note on Y11: Peter W Phillips: History of problems with Vcc. Has link Vcc to VccSense at PPB1 due to intermittent fault, but modules still did not work.
Note on Y23: Peter W Phillips: Zero IPIN during SR1 test, however diode test OK.
Note on Y25: Peter W Phillips: Only one working vcsel (known from Oxford)

## Octant 7
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 7 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
10 | B4 P35 | US15 | 28-15/3 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
3 | B5 P38 | US15 | 29-15/2 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
5 | B5 P39 | US15 | 29-15/2 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
7 | B5 P40 | US15 | 29-15/2 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
6 | B6 P46 | US15 | 29-15/3 | 0 | -6 | ok | 1 | -5 | ok | 2 | -4 | ok | 3 | -3 | ok | 4 | -2 | ok | 5 | -1 | ok | 6 | 1 | ok | 7 | 2 | ok | 8 | 3 | ok | 9 | 4 | ok | 10 | 5 | ok | 11 | 6 | ok
8 | B6 P47 | US15 | 29-15/3 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
9 | B6 P48 | US15 | 29-15/3 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
11 | B6 P49 | US15 | 29-15/3 | 36 | -6 | pass | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
1 | B6 P43 | US15 | 29-15/4 | 12 | -6 | ok | 13 | -5 | ok | 14 | -4 | ok | 15 | -3 | ok | 16 | -2 | ok | 17 | -1 | ok | 18 | 1 | ok | 19 | 2 | ok | 20 | 3 | ok | 21 | 4 | ok | 22 | 5 | ok | 23 | 6 | ok
2 | B6 P44 | US15 | 29-15/4 | 24 | -6 | ok | 25 | -5 | ok | 26 | -4 | ok | 27 | -3 | ok | 28 | -2 | ok | 29 | -1 | ok | 30 | 1 | ok | 31 | 2 | ok | 32 | 3 | ok | 33 | 4 | ok | 34 | 5 | ok | 35 | 6 | ok
4 | B6 P45 | US15 | 29-15/4 | 36 | -6 | ok | 37 | -5 | ok | 38 | -4 | ok | 39 | -3 | ok | 40 | -2 | ok | 41 | -1 | ok | 42 | 1 | ok | 43 | 2 | ok | 44 | 3 | ok | 45 | 4 | ok | 46 | 5 | ok | 47 | 6 | ok
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
21 | B3 P28 | US15 | 28-15/2 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
18 | B3 P27 | US15 | 28-15/2 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
16 | B3 P26 | US15 | 28-15/2 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
22 | B5 P42 | US15 | 28-15/3 | 0 | 6 |  | 1 | 5 |  | 2 | 4 |  | 3 | 3 |  | 4 | 2 |  | 5 | 1 |  | 6 | -1 |  | 7 | -2 |  | 8 | -3 |  | 9 | -4 |  | 10 | -5 |  | 11 | -6
19 | B4 P34 | US15 | 28-15/3 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
14 | B3 P25 | US15 | 29-15/1 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
13 | B4 P31 | US15 | 29-15/1 | 12 | -6 |  | 13 | -5 |  | 14 | -4 |  | 15 | -3 |  | 16 | -2 |  | 17 | -1 |  | 18 | 1 |  | 19 | 2 |  | 20 | 3 |  | 21 | 4 |  | 22 | 5 |  | 23 | 6
15 | B4 P32 | US15 | 29-15/1 | 24 | -6 |  | 25 | -5 |  | 26 | -4 |  | 27 | -3 |  | 28 | -2 |  | 29 | -1 |  | 30 | 1 |  | 31 | 2 |  | 32 | 3 |  | 33 | 4 |  | 34 | 5 |  | 35 | 6
17 | B4 P33 | US15 | 29-15/1 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
20 | B5 P41 | US15 | 29-15/2 | 36 | -6 |  | 37 | -5 |  | 38 | -4 |  | 39 | -3 |  | 40 | -2 |  | 41 | -1 |  | 42 | 1 |  | 43 | 2 |  | 44 | 3 |  | 45 | 4 |  | 46 | 5 |  | 47 | 6
12 | B5 P37 | US15 | 29-15/4 | 0 | -6 |  | 1 | -5 |  | 2 | -4 |  | 3 | -3 |  | 4 | -2 |  | 5 | -1 |  | 6 | 1 |  | 7 | 2 |  | 8 | 3 |  | 9 | 4 |  | 10 | 5 |  | 11 | 6
Note on G10: CableTest: HV backplane resistance static at 42k
Note on G12: CableTest: HV backplane resistance 40k passes cut eventually
Note on G13: CableTest: Some kind of short VddRet to ground? -retested 15-11 - still same fault, then replugged and PASS!
Note on M12: CableTest: HV backplane resistance 40k passes cut eventually
Note on P9: CableTest: HV backplane resistance 60k, fell down to 39k at second test, then static
Note on P22: Peter W Phillips: SR1: high PIN forward voltage (1.35V)
Note on V24: Peter W Phillips: SR1: HV bias connection PPB1-LMT repaired by wire, solder and conductive epoxy
Note on Y9: CableTest: HV backplane resistance 80k fell to 39k at second test
Note on Y13: Peter W Phillips: SR1: Vcc sense repair was made?
Note on Y15: Peter W Phillips: Oxford test: zero pin current, but seems electrical connectivity OK.
Note on AN9: CableTest: HV backplane resistance static at 42k. Only one working vcsel (known from Oxford) but apparently not due to LMT fault!
Note on AN12: CableTest: HV backplane R circa 60k static
Note on AN13: CableTest: Select-VddRet 4.9k. Cut is 5.0k

## Octant 8
ATLAS SCT BARREL CONNECTION TEST LOG SHEET
modules at C side (Z-) [C -> sea -> blue]
| OCTANT 8 |  |  |  | modules at A side (Z+) [A-> hay -> yellow]
| INNER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test
33 | B6 P56 | US15 | 05-14/1 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
32 | B6 P55 | US15 | 05-14/1 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
30 | B6 P54 | US15 | 05-14/1 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
31 | B5 P47 | US15 | 05-14/2 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
29 | B5 P46 | US15 | 05-14/2 | 36 | 6 | ok | 37 | 5 | ok | 38 | 4 | ok | 39 | 3 | ok | 40 | 2 | ok | 41 | 1 | ok | 42 | -1 | ok | 43 | -2 | ok | 44 | -3 | ok | 45 | -4 | ok | 46 | -5 | ok | 47 | -6 | ok
28 | B6 P53 | US15 | 28-15/1 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
26 | B6 P52 | US15 | 28-15/1 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
27 | B5 P45 | US15 | 28-15/1 | 24 | 6 | ok | 25 | 5 | ok | 26 | 4 | ok | 27 | 3 | ok | 28 | 2 | ok | 29 | 1 | ok | 30 | -1 | ok | 31 | -2 | ok | 32 | -3 | ok | 33 | -4 | ok | 34 | -5 | ok | 35 | -6 | ok
24 | B4 P36 | US15 | 28-15/3 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
25 | B6 P51 | US15 | 28-15/4 | 0 | 6 | ok | 1 | 5 | ok | 2 | 4 | ok | 3 | 3 | ok | 4 | 2 | ok | 5 | 1 | ok | 6 | -1 | ok | 7 | -2 | ok | 8 | -3 | ok | 9 | -4 | ok | 10 | -5 | ok | 11 | -6 | ok
23 | B6 P50 | US15 | 28-15/4 | 12 | 6 | ok | 13 | 5 | ok | 14 | 4 | ok | 15 | 3 | ok | 16 | 2 | ok | 17 | 1 | ok | 18 | -1 | ok | 19 | -2 | ok | 20 | -3 | ok | 21 | -4 | ok | 22 | -5 | ok | 23 | -6 | ok
| OUTER | Cavern | Rack | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch | pp | Test | Ch |  | Test | Ch |  | Test | Ch |  | Test | Ch |  | Test | Ch |  | Test | Ch |  | Test
44 | B5 P48 | US15 | 05-14/2 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
43 | B4 P40 | US15 | 05-14/3 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
41 | B4 P39 | US15 | 05-14/3 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
39 | B4 P38 | US15 | 05-14/3 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
42 | B3 P32 | US15 | 05-14/4 | 12 | 6 |  | 13 | 5 |  | 14 | 4 |  | 15 | 3 |  | 16 | 2 |  | 17 | 1 |  | 18 | -1 |  | 19 | -2 |  | 20 | -3 |  | 21 | -4 |  | 22 | -5 |  | 23 | -6
40 | B3 P31 | US15 | 05-14/4 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
38 | B3 P30 | US15 | 05-14/4 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
37 | B4 P37 | US15 | 28-15/1 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
35 | B3 P29 | US15 | 28-15/2 | 0 | 6 |  | 1 | 5 |  | 2 | 4 |  | 3 | 3 |  | 4 | 2 |  | 5 | 1 |  | 6 | -1 |  | 7 | -2 |  | 8 | -3 |  | 9 | -4 |  | 10 | -5 |  | 11 | -6
36 | B5 P44 | US15 | 28-15/4 | 24 | 6 |  | 25 | 5 |  | 26 | 4 |  | 27 | 3 |  | 28 | 2 |  | 29 | 1 |  | 30 | -1 |  | 31 | -2 |  | 32 | -3 |  | 33 | -4 |  | 34 | -5 |  | 35 | -6
34 | B5 P43 | US15 | 28-15/4 | 36 | 6 |  | 37 | 5 |  | 38 | 4 |  | 39 | 3 |  | 40 | 2 |  | 41 | 1 |  | 42 | -1 |  | 43 | -2 |  | 44 | -3 |  | 45 | -4 |  | 46 | -5 |  | 47 | -6
Note on G15: Peter W Phillips: Oxford: VddSense-Vdd connected at PPB1
Note on J8: CableTest: trained well
Note on J12: CableTest: HV trains to 43k
Note on J15: CableTest: after 3S training
Note on M9: CableTest: HV backplane R 54k static
Note on M16: CableTest: 41.4k!!
Note on P8: CableTest: trainedamd stayed then at 41k
Note on P9: CableTest: HV backplane R 42k static
Note on P11: CableTest: HV backplane R 46k static
Note on P13: CableTest: after training 39k
Note on S12: CableTest: 80k, slow training
Note on S13: CableTest: 58k after training
Note on S28: Peter W Phillips: SR1: slightly high R in Vcc loop
Note on V9: CableTest: trained well
Note on V11: CableTest: ok after training
Note on V15: CableTest: 58k
Note on V16: CableTest: Pass after 3S training
Note on Y6: CableTest: trained ok in second try
Note on AE8: CableTest: trained well
Note on AE9: CableTest: HV R stuck at 41k
Note on AE27: Peter W Phillips: Only one working vcsel (known from Oxford)
Note on AH16: CableTest: 42k after 13S training
Note on AH23: Peter W Phillips: SR1: HVRET linked to VccRet Sense at PPB1 due to LMT fault.
Note on AN7: CableTest: trained well
Note on AN9: CableTest: trained first then HV R stuck at 41k
Note on AN22: Peter W Phillips: Oxford test: zero ipin, but diode test OK

## OneCrate
| C |  | 7 |  |  | 15 |  |  | 23 |  |  | 31 |  |  | 39 |  |  | 47 |  | S
| R | [18 empty] | H
S | A |  | 6 |  |  | 14 |  |  | 22 |  |  | 30 |  |  | 38 |  |  | 46 |  | O
C | T | [18 empty] | R
T | E |  | 5 |  |  | 13 |  |  | 21 |  |  | 29 |  |  | 37 |  |  | 45 |  | T
I
I | C |  | 4 |  |  | 12 |  |  | 20 |  |  | 28 |  |  | 36 |  |  | 44 |  | N
N | O | [18 empty] | G
T | N | 3 | 3 | 7 | 11 | 11 | 15 | 19 | 19 | 23 | 27 | 27 | 31 | 35 | 35 | 39 | 43 | 43 | 47
E | T | [18 empty] | C
R | R | 2 | 2 | 6 | 10 | 10 | 14 | 18 | 18 | 22 | 26 | 26 | 30 | 34 | 34 | 38 | 42 | 42 | 46 | A
L | O | [18 empty] | R
O | L | 1 | 1 | 5 | 9 | 9 | 13 | 17 | 17 | 21 | 25 | 25 | 29 | 33 | 33 | 37 | 41 | 41 | 45 | D
C | L
K | E | 0 | 0 | 4 | 8 | 8 | 12 | 16 | 16 | 20 | 24 | 24 | 28 | 32 | 32 | 36 | 40 | 40 | 44
R
|  | LV | HV | LV | LV | HV | LV | LV | HV | LV | LV | HV | LV | LV | HV | LV | LV | HV | LV

## Crates
ATLSCTUSBPS1 | Location | 29-15 / 1 | 29-15 / 2 | 29-15 / 3 | 29-15 / 4 | 28-15 / 1 | 28-15 / 2 | 28-15 / 3 | 28-15 / 4 | 27-15 / 1 | 24-15 / 1 | 24-15 / 2
| ELMB | 0X01 | 0X02 | 0X03 | 0X04 | 0X05 | 0X06 | 0X07 | 0X08 | 0X09 | 0X15 | 0X16 | A/B
| CRATE | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 20 | 21
| 00-11 | 3025 | 5038 | 6046 | 5037 | 6053 | 3029 | 5042 | 6051 | 6014 | 3006 | 4006
| 12-23 | 4031 | 5039 | 6047 | 6043 | 6052 | 3028 | 4036 | 6050 | 5012 | 3007 | 4007
| 24-35 | 4032 | 5040 | 6048 | 6044 | 5045 | 3027 | 4035 | 5044 | 5011 | 3008 | 4008
| 36-47 | 4033 | 5041 | 6049 | 6045 | 4037 | 3026 | 4034 | 5043 | 5010 | 4010 | 4009
ATLSCTUSBPS2 | Location | 24-15 / 3 | 24-15 / 4 | 23-15 / 1 | 23-15 / 2 | 23-15 / 3 | 23-15 / 4 | 5-14 / 1 | 5-14 / 2 | 5-14 / 3 | 5-14 / 4 | 12-15 / 1
| ELMB | 0X17 | 0X18 | 0X19 | 0X1A | 0X1B | 0X1C | 0X1D | 0X1E | 0X1F | 0X20 | 0X29 | E/H
| CRATE | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 40
| 00-11 | 5006 | 6010 | 3002 | 4002 | 5002 | 6006 | 6001 | 5001 | 4001 | 3001 | 6004
| 12-23 | 5007 | 6011 | 3003 | 4003 | 5003 | 6007 | 6056 | 5048 | 4040 | 3032 | 6005
| 24-35 | 5008 | 6012 | 3004 | 4004 | 5004 | 6008 | 6055 | 5047 | 4039 | 3031 | 6002
| 36-47 | 5009 | 6013 | 3005 | 4005 | 5005 | 6009 | 6054 | 5046 | 4038 | 3030 | 6003
ATLSCTUSEPS1 | Location | 27-15 / 2 | 27-15 / 3 | 27-15 / 4 | 26-15 / 1 | 36-15 / 2 | 26-15 / 3 | 26-15 / 4 | 25-15 / 1 | 25-15 / 2 | 25-15 / 3 | 25-15 / 4
| ELMB | 0X0A | 0X0B | 0X0C | 0X0D | 0X0E | 0X0F | 0X10 | 0X11 | 0X12 | 0X13 | 0X14 | C/D
| CRATE | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19
ATLSCTUSEPS2 | Location | 11-15 / 1 | 11-15 / 2 | 12-15 / 2 | 12-15 / 3 | 12-15 / 4 | 10-15 / 1 | 10-15 / 2 | 10-15 / 3 | 10-15 / 4 | 11-15 / 3 | 11-15 / 4
| ELMB | 0X25 | 0X26 | 0X2A | 0X2B | 0X2C | 0X21 | 0X22 | 0X23 | 0X24 | 0X27 | 0X28 | F/G
| CRATE | 36 | 37 | 41 | 42 | 43 | 32 | 33 | 34 | 35 | 38 | 39
|  |  | Octant 1 | Octant 2 | Octant 3 | Octant 4 | Octant 5 | Octant 6 | Octant 7 | Octant 8
ATLSCTUSABPS1 | Location | 22-14 / 1 | 22-14 / 2 | 22-14 / 3 | 26-14 / 1 | 28-14 / 1 | 26-11 / 1 | 26-11 / 2 | 25-11 / 1 | 25-11 / 2 | 25-11 / 3 | 25-11 / 4
| ELMB | 0X01 | 0X02 | 0X03 | 0X09 | 0X11 | 0X25 | 0X26 | 0X29 | 0X2A | 0X2B | 0X2C | B/I
| CRATE | 44 | 45 | 46 | 52 | 60 | 80 | 81 | 84 | 85 | 86 | 87
| 00-11 | 5014 | 6018 | 5013 | 5033 | 6034 | 3013 | 4016 | 3022 | 4026 | 5034 | 6038
| 12-23 | 5015 | 6019 | 6015 | 5032 | 6035 | 3012 | 4015 | 3023 | 4027 | 5035 | 6039
| 24-35 | 5016 | 6020 | 6016 | 5031 | 6036 | 3011 | 4014 | 3024 | 4028 | 5036 | 6040
| 36-47 | 5017 | 6021 | 6017 | 5030 | 6037 | 3010 | 6023 | 4030 | 4029 | 6042 | 6041
ATLSCTUSABPS2 | Location | 28-14 / 2 | 28-14 / 3 | 28-14 / 4 | 29-14 / 1 | 29-14 / 2 | 29-11 / 1 | 29-11 / 3 | 28-11 / 3 | 28-11 / 4 | 26-11 / 3 | 26-11 / 4
| ELMB | 0X12 | 0X13 | 0X14 | 0X15 | 0X16 | 0X19 | 0X1B | 0X1F | 0X20 | 0X27 | 0X28 | F/H
| CRATE | 61 | 62 | 63 | 64 | 65 | 68 | 70 | 74 | 75 | 82 | 83
| 00-11 | 4023 | 3019 | 6033 | 6030 | 5025 | 5022 | 3014 | 6024 | 6026 | 5020 | 4013
| 12-23 | 3018 | 4024 | 5027 | 6031 | 4021 | 5023 | 3015 | 6025 | 6027 | 5019 | 4012
| 24-35 | 3020 | 6032 | 5028 | 5026 | 3017 | 4019 | 3016 | 5021 | 6028 | 5018 | 4011
| 36-47 | 3021 | 4025 | 5029 | 4022 | 6029 | 4020 | 5024 | 4017 | 4018 | 6022 | 3009
ATLSCTUSAEPS1 | Location | 22-14 / 4 | 23-14 / 1 | 23-14 / 2 | 23-14 / 3 | 23-14 / 4 | 27-14 / 3 | 27-14 / 4 | 29-14 / 3 | 29-14 / 4 | 29-11 / 2 | 29-11 / 4
| ELMB | 0X04 | 0X05 | 0X06 | 0X07 | 0X08 | 0X0F | 0X10 | 0X17 | 0X18 | 0X1A | 0X1C | C/E
| CRATE | 47 | 48 | 49 | 50 | 51 | 58 | 59 | 66 | 67 | 69 | 71
ATLSCTUSAEPS2 | Location | 26-14 / 2 | 26-14 / 3 | 26-14 / 4 | 27-14 / 1 | 27-14 / 2 | 28-11 / 1 | 28-11 / 2 | 27-11 / 1 | 27-11 / 2 | 27-11 / 3 | 27-11 / 4
| ELMB | 0X0A | 0X0B | 0X0C | 0X0D | 0X0E | 0X1D | 0X1E | 0X21 | 0X22 | 0X23 | 0X24 | D/G
| CRATE | 53 | 54 | 55 | 56 | 57 | 72 | 73 | 76 | 77 | 78 | 79
|  | 53: "Upside Down" racks in Italics. Crate numbers still count from 1 (top) to 4 (bottom)!; 79: Last Update | 2006-11-03

## US15
US15
Branch | Description | PC | Project | System | CAN card | CAN port | #crates
A | PS crate controller branch B (ELMB custom) | PCATLSCTUSBPS1 | ATLSCTUSBPS1 | 20 | 0 | 0 | 6
B | PS crate controller branch B (ELMB custom) | PCATLSCTUSBPS1 | ATLSCTUSBPS1 | 20 | 0 | 1 | 5
C | PS crate controller branch EC (ELMB custom) | PCATLSCTUSEPS1 | ATLSCTUSEPS1 | 22 | 0 | 0 | 6
D | PS crate controller branch EC (ELMB custom) | PCATLSCTUSEPS1 | ATLSCTUSEPS1 | 22 | 0 | 1 | 5
E | PS crate controller branch B (ELMB custom) | PCATLSCTUSBPS2 | ATLSCTUSBPS2 | 21 | 0 | 0 | 6
F | PS crate controller branch EC (ELMB custom) | PCATLSCTUSEPS2 | ATLSCTUSEPS2 | 23 | 0 | 0 | 5
G | PS crate controller branch EC (ELMB custom) | PCATLSCTUSEPS2 | ATLSCTUSEPS2 | 23 | 0 | 1 | 6
H | PS crate controller branch B (ELMB custom) | PCATLSCTUSBPS2 | ATLSCTUSBPS2 | 21 | 0 | 1 | 5
I | Power Pack Monitor (ELMB) | PCATLSCTUSENV | ATLSCTUSENV | 28 | 0 | 0
J | Power Pack Monitor (ELMB) | PCATLSCTUSENV | ATLSCTUSENV | 28 | 0 | 1
K | BBIM crate (Y.34-05.X.8) branch (ELMB) | PCATLSCTUSENV | ATLSCTUSENV | 28 | 0 | 2
L | BBIM crate (Y.29.04.X.1) branch (ELMB) | PCATLSCTUSENV | ATLSCTUSENV | 28 | 0 | 3
Address | Crate Location | Barrel or EndCap | CAN bus
0x01 | 29-15.S2 - 1 | Barrel | A
0x02 | 29-15.S2 - 2 | Barrel | A
0x03 | 29-15.S2 - 3 | Barrel | A
0x04 | 29-15.S2 - 4 | Barrel | A
0x05 | 28-15.S2 - 1 | Barrel | A
0x06 | 28-15.S2 - 2 | Barrel | A
0x07 | 28-15.S2 - 3 | Barrel | B
0x08 | 28-15.S2 - 4 | Barrel | B
0x09 | 27-15.S2 - 1 | Barrel | B
0x0A | 27-15.S2 - 2 | Endcap | C
0x0B | 27-15.S2 - 3 | Endcap | C
0x0C | 27-15.S2 - 4 | Endcap | C
0x0D | 26-15.S2 - 1 | Endcap | C
0x0E | 26-15.S2 - 2 | Endcap | C
0x0F | 26-15.S2 - 3 | Endcap | C
0x10 | 26-15.S2 - 4 | Endcap | D
0x11 | 25-15.S2 - 1 | Endcap | D
0x12 | 25-15.S2 - 2 | Endcap | D
0x13 | 25-15.S2 - 3 | Endcap | D
0x14 | 25-15.S2 - 4 | Endcap | D
0x15 | 24-15.S2 - 1 | Barrel | B
0x16 | 24-15.S2 - 2 | Barrel | B
0x17 | 24-15.S2 - 3 | Barrel | E
0x18 | 24-15.S2 - 4 | Barrel | E
0x19 | 23-15.S2 - 1 | Barrel | E
0x1A | 23-15.S2 - 2 | Barrel | E
0x1B | 23-15.S2 - 3 | Barrel | E
0x1C | 23-15.S2 - 4 | Barrel | E
0x1D | 5-14.S2 - 1 | Barrel | H
0x1E | 5-14.S2 - 2 | Barrel | H
0x1F | 5-14.S2 - 3 | Barrel | H
0x20 | 5-14.S2 - 4 | Barrel | H
0x21 | 10-15.S2 - 1 | Endcap | G
0x22 | 10-15.S2 - 2 | Endcap | G
0x23 | 10-15.S2 - 3 | Endcap | G
0x24 | 10-15.S2 - 4 | Endcap | G
0x25 | 11-15.S2 - 1 | Endcap | F
0x26 | 11-15.S2 - 2 | Endcap | F
0x27 | 11-15.S2 - 3 | Endcap | G
0x28 | 11-15.S2 - 4 | Endcap | G
0x29 | 12-15.S2 - 1 | Barrel | H
0x2A | 12-15.S2 - 2 | Endcap | F
0x2B | 12-15.S2 - 3 | Endcap | F
0x2C | 12-15.S2 - 4 | Endcap | F |  | Last Update | 2006-10-30

## USA15
USA15
Branch | Description | PC | Project | System | CAN card | CAN port | #crates
A | ROD crate fan tray branch (WIENER) | PCATLSCTUSAENV | ATLSCTUSAENV | 29 | 0 | 0
B | PS crate controller branch B (ELMB custom) | PCATLSCTUSABPS1 | ATLSCTUSABPS1 | 24 | 0 | 0 | 5
C | PS crate controller branch EC (ELMB custom) | PCATLSCTUSAEPS1 | ATLSCTUSAEPS1 | 26 | 0 | 0 | 5
D | PS crate controller branch EC (ELMB custom) | PCATLSCTUSAEPS2 | ATLSCTUSAEPS2 | 27 | 0 | 0 | 5
E | PS crate controller branch EC (ELMB custom) | PCATLSCTUSAEPS1 | ATLSCTUSAEPS1 | 26 | 0 | 1 | 6
F | PS crate controller branch B (ELMB custom) | PCATLSCTUSABPS2 | ATLSCTUSABPS2 | 25 | 0 | 0 | 5
G | PS crate controller branch EC (ELMB custom) | PCATLSCTUSAEPS2 | ATLSCTUSAEPS2 | 27 | 0 | 1 | 6
H | PS crate controller branch B (ELMB custom) | PCATLSCTUSABPS2 | ATLSCTUSABPS2 | 25 | 0 | 1 | 6
I | PS crate controller branch B (ELMB custom) | PCATLSCTUSABPS1 | ATLSCTUSABPS1 | 24 | 0 | 1 | 6
J | Power Pack Monitor (ELMB) | PCATLSCTUSAENV | ATLSCTUSAENV | 29 | 0 | 1
K | BOC Monitoring (ELMB) | PCATLSCTUSAENV | ATLSCTUSAENV | 29 | 1 | 0
L | BBIM crate (Y.29-23.X.8) branch (ELMB) | PCATLSCTUSAENV | ATLSCTUSAENV | 29 | 1 | 1
M | BBIM crate (Y.34-23.X.0) branch (ELMB) | PCATLSCTUSAENV | ATLSCTUSAENV | 29 | 1 | 2
Address | Crate Location | Barrel or EndCap | CAN bus
0x01 | 22-14A2-1 | Barrel | B
0x02 | 22-14A2-2 | Barrel | B
0x03 | 22-14A2-3 | Barrel | B
0x04 | 22-14A2-4 | Endcap | C
0x05 | 23-14A2-1 | Endcap | C
0x06 | 23-14A2-2 | Endcap | C
0x07 | 23-14A2-3 | Endcap | C
0x08 | 23-14A2-4 | Endcap | C
0x09 | 26-14A2-1 | Barrel | B
0x0A | 26-14A2-2 | Endcap | D
0x0B | 26-14A2-3 | Endcap | D
0x0C | 26-14A2-4 | Endcap | D
0x0D | 27-14A2-1 | Endcap | D
0x0E | 27-14A2-2 | Endcap | D
0x0F | 27-14A2-3 | Endcap | E
0x10 | 27-14A2-4 | Endcap | E
0x11 | 28-14A2-1 | Barrel | B
0x12 | 28-14A2-2 | Barrel | F
0x13 | 28-14A2-3 | Barrel | F
0x14 | 28-14A2-4 | Barrel | F
0x15 | 29-14A2-1 | Barrel | F
0x16 | 29-14A2-2 | Barrel | F
0x17 | 29-14A2-3 | Endcap | E
0x18 | 29-14A2-4 | Endcap | E
0x19 | 29-11A2-1 | Barrel | H
0x1A | 29-11A2-2 | Endcap | E
0x1B | 29-11A2-3 | Barrel | H
0x1C | 29-11A2-4 | Endcap | E
0x1D | 28-11A2-1 | Endcap | G
0x1E | 28-11A2-2 | Endcap | G
0x1F | 28-11A2-3 | Barrel | H
0x20 | 28-11A2-4 | Barrel | H
0x21 | 27-11A2-1 | Endcap | G
0x22 | 27-11A2-2 | Endcap | G
0x23 | 27-11A2-3 | Endcap | G
0x24 | 27-11A2-4 | Endcap | G
0x25 | 26-11A2-1 | Barrel | H
0x26 | 26-11A2-2 | Barrel | H
0x27 | 26-11A2-3 | Barrel | I
0x28 | 26-11A2-4 | Barrel | I
0x29 | 25-11A2-1 | Barrel | I
0x2A | 25-11A2-2 | Barrel | I
0x2B | 25-11A2-3 | Barrel | I
0x2C | 25-11A2-4 | Barrel | I |  | Last Update | 2006-10-30

## Sheet1
B6 P49 C side module 6 | [5 empty] | Shield | VddRet | VccRSen | VddRSen | T1 | T2 | VPIN | VVCSEL | SELECT | HVRet | RESET | Vdd | Vcc | HVBias
[6 empty] | to VddRet | to VccRet | to VccRet | to VddRet | to VddRet | to VddRet | to VddRet | to VddRet | to VddRet | to VddRet | to Vdd | to VddSen | to VccSen | to HvRet
2006-11-14 | 17:38 | US15 | 29-15.S2 / 3 | Channel36 | FAIL | 19600 | 744 | 11.4 | 19700 | 26300 | 26300 | -0.315 | 149000 | 23500 | 19700 | 950 | 11.1 | 11.7 | 31400 | battery off
2006-11-14 | 17:39 | US15 | 29-15.S2 / 3 | Channel36 | FAIL | 19700 | 744 | 11.4 | 19700 | 26300 | 26300 | -0.306 | 150000 | 23600 | 19800 | 950 | 11.1 | 11.7 | 31400 | battery off
2006-11-14 | 17:39 | US15 | 29-15.S2 / 3 | Channel36 | FAIL | -89900 | 747 | 11.4 | -89800 | -77600 | -77500 | -1.61 | 47700 | -84100 | -88900 | 950 | 11.1 | 11.7 | 31400 | battery on
2006-11-14 | 17:40 | US15 | 29-15.S2 / 3 | Channel36 | FAIL | -89900 | 747 | 11.4 | -89800 | -77600 | -77500 | -1.61 | 47700 | -84100 | -88900 | 950 | 11.1 | 11.7 | 31400 | battery on